| Settings for EPR |  |  |
| --- | --- | --- |
| Instrument | Sweep Mode Settings | Mod/Locking Mode Settings |
| Function Generator (RIGOL DG1022) Ch 1 | Front Panel Buttons Local: Sweep Waveform: Sine Sweep Settings Type: Linear Sweep Time: 100ms Return Time: 0.0ms Center: Transition freq. Span: ~1Mhz Sine Settings Period: Same as sweep time RF Amplitude: 2V peak to peak Utilities Settings Impedence: 50 | Front Panel Buttons Local: Mod Waveform: Sine Mod Settings Source: External Size: 10kHz/±5V Center: transition freq. Sine Settings Frequency: 100Hz RF Amplitude: 2V peak to peak Utilities Settings Impedence: 50 |
| Function Generator (RIGOL DG1022) Sync 1 | Waveform: Square (default) Frequency: 10Hz (set by Sweep Time) Amplitude: 1.7V peak to peak Impedance: 50 | Waveform: Square (default) Frequency: 100Hz Amplitude: 5V peak to peak Impedance: 50 |
| Function Generator (RIGOL DG1022) Ch 2 | Disabled | Front Panel Buttons Waveform: Square Square Wave Settings Frequency: 200 Hz Amplitude: 3V peak to peak Utilities Settings Impedence: 50 (utility > channel set > output set > impd) |
| Function Generator (RIGOL DG1022) Sync 2 | Disabled | Waveform: Square (default shape) Frequency: 200Hz Amplitude: 1.7V peak to peak Impedance: 50 |
| Lock-in Amplifier (SR530) | Unlocked | Signal Filters Bandpass: In Line: Out Line x2: Out Other Settings Sensitivity: 10mV Trigger: Square Time Constant: 100ms Phase: 0 |

# EPR Setup Instructions

## Function generator

- Set function generator to the appropriate settings as detailed in settings chart

## Probe Beam

- Turn on the probe beam to ~60mA (~0.250 on LDT)) , well detuned from D2 resonance. (You will want to run this on low power for noise reduction. The large detuning is so that the probe beam doesn’t get absorbed and disrupt the measurement.)

- Set up probe beam to pass through the cell in the transverse direction

- Set Glan Thompson Prism to 45 so that 45 degree linearly polarized light is what passes through the cell.

- Verify that the + and – channels of the balance photodiode are balanced.

## Pump Beam

- Turn on pump beam(s) to ~200mA (~0.800 on LTD).

- Set the pump beam to the D1 resonance.

  - Find resonance using the monitor. You will see florescence throughout the cell, with a min florescence at the front of the cell corresponds resonance for D1.

## EPR Coil

- Tune (or check the tune of) the EPR coil using the network analyzer. The coil’s resonance should be off the big discontinuity in phase as much as possible.

  - With coil tuned such that RF broadening actually happens, at 80C the amount of attenuation used is ~10